# Sifab AS — Request for Quotation

# Seraphin Calibration Can — Snorre A Small Volume Prover

| Field | Value |
| --- | --- |
| Date | 2026-02-24 |
| From | Sifab AS (Tom Sverre Falch / Sondre Falch) |
| To | Pemberton Fabricators, Inc. |
| Project | Snorre A — Small Volume Prover (Equinor, Norwegian North Sea) |
| Sifab Reference | SP-01415 |
| End Client | Equinor — Snorre A platform |

____________________________________________________________

## 1. Background

Sifab AS is supplying a Honeywell Small Volume Prover (SVP085) for Equinor's Snorre A platform in the Norwegian North Sea. We require a Seraphin calibration can for water draw testing at the factory (FAT) and on the platform (SAT).

Sifab has previously purchased two Seraphin cans from Pemberton for our NOA/Valhall prover projects:

151.4 Liter — Series M, S/N 24-67715, Cert #24228 (Feb 2024)

75.7 Liter — Series M, S/N 24-67716, Cert #24231 (Feb 2024)

We are very satisfied with the quality and wish to order another can for the Snorre A project.

____________________________________________________________

## 2. Item Required

| # | Item | Specification |
| --- | --- | --- |
| 1 | Seraphin Calibration Can (Series M Prover) | Volume to be confirmed — likely larger than 151.4L based on SVP085 prover displacement. Please advise recommended volume for a Honeywell SVP085 prover. |

### 2.1 Technical Requirements

| Parameter | Requirement |
| --- | --- |
| Material | Type 316 Stainless Steel (all wetted parts) |
| Calibration standard | NIST Handbook 105-3 and API MPMS Chapter 4 |
| Calibration type | "To Deliver" at 15°C reference temperature |
| Scale accuracy | ≤0.02% |
| Scale type | "High Accuracy" neck (≥1.5% volume) |
| Drain valves | (2) SS ball drain valves at 7-degree decline |
| Thermometer well | (1) SS thermometer well at center |
| Level indicator | T-type with aluminum graduated scale |
| Data plate | Engraved brass |
| Gauge glass fittings | SS |
| Lifting loops | (2) for overhead crane handling |
| Leveling feet | (3) adjustable |
| Calibration laboratory | NVLAP-accredited, NIST-traceable |

### 2.2 Certification Requirements

The Seraphin can must be certified by Justervesenet (the Norwegian National Metrology Service) for use in fiscal metering calibration. Please advise:

Can Pemberton arrange Justervesenet certification directly?

Or should Sifab coordinate this separately?

Estimated lead time for certification?

### 2.3 Operating Environment

| Parameter | Value |
| --- | --- |
| Location | Offshore platform (Snorre A, Norwegian North Sea) |
| Storage | In SS316L protective cabinet (manufactured separately by Sifab) |
| Fluid for water draw | Clean water |
| Ambient temperature (storage) | -8°C to +40°C |
| Use frequency | Periodic calibration (typically annual) |

____________________________________________________________

## 3. Quotation Format Requested

Please provide:

Seraphin can unit price — with all standard features listed above

Recommended volume for a Honeywell SVP085 prover

Calibration certificate — NVLAP laboratory, NIST-traceable

Justervesenet certification — if available, price and lead time

Shipping — FOB Rancocas, NJ or CIF Sandnes, Norway (quote both if possible)

Delivery time — weeks from order

Validity of quotation

____________________________________________________________

## 4. Delivery

| Parameter | Requirement |
| --- | --- |
| Delivery address (option 1) | Honeywell/TruStop, 1725 W. 10th Pl., Tempe, AZ 85281 (for FAT use) |
| Delivery address (option 2) | Sifab AS, Bedriftsveien 24, 4313 Sandnes, Norway |
| Packing | Suitable for international sea/air freight |

Please quote both delivery options.

____________________________________________________________

## 5. Previous Order Reference

For your records, our previous order:

Pemberton S/N 24-67715 — 151.4L Series M, calibrated Feb 2024, Cert #24228

Pemberton S/N 24-67716 — 75.7L Series M, calibrated Feb 2024, Cert #24231

Calibration by Pennsylvania Standards Laboratory (NVLAP Lab Code 200869-0)

Contact: Ron Gibson

____________________________________________________________

## 6. Contacts

| Role | Name | Phone | Email |
| --- | --- | --- | --- |
| Sales | Sondre Falch | +47 900 29 588 | sondre.falch@sifab.no |
| Technical / PM | Tom Sverre Falch | +47 416 28 408 | tom.falch@sifab.no |

Sifab AS

Bedriftsveien 24, 4313 Sandnes, Norway

www.sifab.no