## Checklist
Prover Checklist EQUINOR SNORRE
Customer Name |  | Sifab AS for Guidant Measurement Solutions (Equinor Snorre A)
Customer PO:
No | Description of Discrepancy | Yes/No | Comments | NOOTES
Entire prover project HON standard documentation is approved.
All documentation shared with HON can be dischared as long as we make sure the punchlist items and TQ's are handeled.
| Lead time flowtube |  | Entire delivery time SVP-O25 - 28 wks. RFQ bid deadline 4 March 2026. Flowtube from US stock (O25, 316SS). CL600 RTJ connections per RFQ. | Since the lead time request will be 28 weeks, and critical part of lead time is the subdelivery of the flow tube: We are allowed to use flowtube from stock US: As long as it's O25 (316ss). Also 300# prc connections are allowed
1 | Are there any AVL requirements; if yes, provide AVL | No | No information in specs
2 | Motor must have direction arrow | Yes |  | Motor can be HON standard applied motor
3 | Motor must have orange end cover | Yes |  | MUNSELL 5 G 8/4 painted.
4 | Motor, JB & control panel must have NR10warning labels | Yes
5 | Are there cable tray mounting requirements? (if yes, proivide requirements in comments) | Yes | SS316L material All cable trays shall be applied with Polyurethane coat before installation to prevent corrosion. All cable trays exposed to direct sunlight and vertical trays shall be provided with cable tray covers. These covers shall be secured with SS 316 clamps using a nut and bolt method and shall also be wrapped with PVC-insulated SS 316 straps. | Protective varnish on cable trays
6 | Cable ties shall be SS with PVC/nylon coating | Yes | Cable ties shall be of PVC-coated SS 316.
7 | cables shall have cable numbers installed at each end | Yes | All cables shall have a SS 316 cable tag with black engraved letter affixed to the cable identifying the cable number at each penetration of a building, enclosure, junction box, panels, etc
8 | Earth strap installed on the prover shall be Green/yellow | Yes
9 | Equipment earth cable shall be 6mm | Yes
10 | Cables shall be LSZH type | Yes
11 | manufacturer nameplate shall be dual language (If yes, provide language in comments) | Yes | Caution Plate: English and Norwegian language (International recognized visual signs shall be utilized). Shall be provided withSS 316 black engraved nameplates permanently attached with SS 316 screws or rivets. The name plate shall include the instrument/equipment tag and description. The tag plate for action items, i.e. E-Stop PB should be provided with service description. Instrument/equipment tagging shall be as per the Company’s tag numbering procedure. Ex label shall be provided with SS 316 plate with SS 316 screws or rivets. Sticker is not acceptable except for glass reinforced plastic JB where silver foiled sticker is to be attached inside and outside of the box.
12 | Control panel door must have earth cable | Yes
13 | Stuffing glands into the bottom of the NEMA 4X box must have a serrated/star washer installed | Yes | IP66 rating required per RFQ section 1.2 (not NEMA 4X)
14 | Adaptors must be Metric to NPT for the cable glands | No | Gland size shall be metric. In general, double compression type gland shall be used for all cables. All cable gland and adaptors will be made of nickel-plated brass or SS 316L Cable glands shall be dual-certified Ex “d” and Ex “e,” IP66, nickel-plated brass, double compression-type.
15 | Minimum allowed size for power cables is 2.5mm | Yes | Power: 2.5 mm2. Control: 1.5 mm2. Instrument: 0.75 mm2. C.T. Secondary: 6 mm2
16 | Motor power cable into the power control must be Brown, Black, Grey | No | Core Identification for 0.6/1 kV Power Cables - The individual cores shall be identified as follows: • 1 core : Grey (Off-white). • 3C + E : Grey (Off-white) + Black + Red + Yellow/Green. • Multi-Core + E and above : White number on black insulation. • Conductors shall be color coded in accordance with BS 7671-514. • Multi conductor cables, which are not provided with color coded conductor, shall have colored, heat shrinkable tubing in accordance with BS 7671-514 installed over the exposed insulation. TBC with client
17 | All cable cores must be terminated | Yes | No information in specs, but we can keep
18 | SVP controller box, motor controller, and terminal box must have tag installed with service description and TAG ID's (if yes, customer must provide all tag ID's prior to kick off meeting) | Yes | All instruments, JBs, cabinets, panels, detectors, pushbuttons, and ancillary equipment shall be provided with SS 316 black engraved nameplates permanently attached with SS 316 screws or rivets. The name plate shall include the instrument/equipment tag and description. The tag plate for action items, i.e. E-Stop PB should be provided with service description. Instrument/equipment tagging shall be as per the Company’s tag numbering procedure. Ex label shall be provided with SS 316 plate with SS 316 screws or rivets. Sticker is not acceptable except for glass reinforced plastic JB where silver foiled sticker is to be attached inside and outside of the box.
| Confirm compliance with ABS quality requirements. Only the costs of the ABS inspector will be covered by SiFAB? |  |  | Copy from OFFSHORE: ISO9001 certified materials for pressurized wetted parts
19 | NACE MR 0175 (2003) or latest/ ISO 15156 requirements | Yes | NACE MR0175 / ISO 15156-3 for H2S and sand particles. Equinor TR2000 BD20X material spec. Gravimetric calibration witnessed by Buyer, Contractor, End Client and Norwegian Authorities. Justervesenet certified Seraphin can required.
20 | Material of construction Flow tube material - SS316/316L NACE Flow tube material - SS316/316L NACE Seal materials - Filled PTFE Seal materials - Filled PTFE Juntion Box - SS316 Weted parts - Dual Certified 316/316L | Yes
21 | Zone 1, Group IIA, Temperature Class T3, IP66 minimum, per ATEX 2014/34/EU and IEC-60079. RFQ specifies Zone 1 IIA T3 as minimum. | Yes
22 | ±0.1% accuracy ±0.02% repeatability | Yes